

Risk Assessment


	Background Information - EXAMPLE
	Date: 

Name of Assessor: 

	Describe the product and application
	The Departmental Confocal Facility contains four lasers connected to the Leica SP and the LSR UltraVIEW confocal heads, which are in turn connected to the Leica DMR (Upright) and Leica DM/IRBE (Inverted) microscopes. The facility is used for the fluorescent imaging of biological specimens.



	Describe the Laser
	The Leica SP confocal head is attached to three lasers:

1. HeNe Laser (JDS Uniphase, Type 1153P), Class 3B, continuous wave, 15mW maximum power (<5 at focal plane), emitted wavelengths 633 nm

2. Argon Ion Laser (JDS Uniphase, Type 2211-65ML), Class 3B, continuous wave, 100mW maximum power (<10 at focal plane), emitted wavelengths 458, 476, 488 and 514 nm

3. Krypton Ion Laser (Melles Griot, Type 643-B01), Class 3B, continuous wave, 25mW maximum power (<5 at focal plane), emitted wavelengths 568nm
The LSR Ultraview head is attached to one laser:

1. Argon-Krypton Ion Laser (Omnichrome 643-LIFE-A01), Class 3B, continuous wave, 80 mW maximum power, emitted wavelengths 488 & 568 nm 

	Describe the Beam Delivery System
	The laser beams are delivered by a closed fibre optic system to the confocal head. In the Leica machine the laser light is merged in the laser merge unit then passed through an AOTF and into the optical system of the microscope. In the UltraVIEW the laser light enters the confocal head and  passes through a series of filters and a version of a Nipkov Disk.


	Describe the Laser Process
	Laser light is applied, through the microscope objective, to the sample. In the Leica the laser beam scans the sample; in the Ultraview a spinning disk directs the beam. Emitted fluorescent light, from appropriate dyes or proteins, is collected through the objective, and detected by PMTs (Leica) or a CCD camera (Ultraview) in the confocal head.


	Describe the Environment
	The system is in a dedicated, purpose-refurbished room, T6. Access to the room is controlled by swipe card. The room itself is enclosed and air conditioned. It contains two microscopes, two confocal apparatus, 4 lasers, two PCs and miscellaneous small equipment.


	Who uses the product or could affect its operation?
	Authorised users (see appendix A for list of authorised users), maintenance staff (see appendix B for list of maintenance staff), service engineers 

	Underline the part(s) of the life cycle of interest
	Planning, Design, Manufacture, Testing, Transport, Installation, Commissioning, Normal Operation, Maintenance, Servicing, Modification, Decommissioning,  Disposal


EXAMPLE – Confocal microscope system 

	Assessment Number:                                                Assessment Date: 

Assessed by:                                                             Review Date:
	Activity/Facility Assessed: Confocal Facility

Location: 

	STEP 1
	STEP 2
	STEP 3

	List significant hazards


	List groups of people who are at risk
	List existing controls
	Are these controls OK?
	What is the risk factor from these hazards?
	Actions Required 

(See over)

	The Laser:

Electric shock from high voltage power supply

Cooling system 

Explosion of mercury arc lamp

Equipment interconnections -Trip hazard and possible damage to equipment (e.g.fibre breakage)

	All


	Lasers are fully contained in cabinets. Only qualified service personnel are permitted to open these boxes.

Instructions to this effect are contained in the local rules and are part of pre-use training required for authorisation.

Laser cabinets are actively ventilated to external source. Integrity of system is checked by maintenance technician.

System is serviced regularly (on contract)

Lamp is changed regularly after ? hours by technician. Housing of lamp is designed to contain explosion. Adequate ventilation and contingency plan devised in case of emergency (set out in local rules)

All optical guides and other equipment interconnections are kept out of the way.


	Yes

Yes

Yes

Yes
	Low

Low

Low

Low
	Add to inspection checklist

Add to inspection checklist

Add to inspection checklist

	Beam Delivery:

Exposure to laser radiation, Class 3B, UV-A and visible wavelengths.

UV-A – Short term – burns to skin

Long term – premature cataracts in lens of eye. Associated with malignant melanoma in skin

Visible – retinal hazard
Laser beam is divergent when objectives in place 

Exposure to collimated Class 3B laser beam if epi-illuminator head tilted back and objective removed 

	All

Maintenance staff


	Beam delivery

1. Optical fibres encased in flexible metal coverings and are attached by screws on fixings at all points. They are routed behind the equipment. Only qualified service personnel are permitted to disconnect this equipment.

2. Beam contained within microscope optics – interlocked to prevent viewing laser through microscope optics.

3. Free air from objective to sample – ? cm gap. Users instructed not to insert reflective material into beam path.

Users may not alter the UV light source. Exposure is limited at the sample. Users are advised to block the UV light source by switching off or using filter when changing samples. Avoid exposure to UV as much as possible.

Instructions contained in the local rules and are part of pre use training required for authorisation.

On some models interlock stops laser beam when head tilted back. Maintenance staff trained by service personnel on how to maintain eqpt safely. Attended laser safety course.
	Yes

Yes


	Low

Low
	Add to inspection checklist

(Check whether this applies to your system)



	The Laser Process:

Reflections of laser radiation from sample and metalwork
Chemical hazards from samples

Biological hazards from samples

Contamination 
	Users

Users

Users

All
	Light intensity at the process site is reduced. During normal use, users are exposed to diffuse scatter from the sample and diffuse scatter is safe to view (Class 3B only). Direct viewing of the beam is not possible due to the design, and exposure to specular reflections is prevented by the following: Extra protection by a light shield of appropriate filter material provided by the manufacturer.

User training and the local rules emphasise shutting off of lasers for changing samples and avoiding reflections of laser light from samples. 

Users are not permitted to change the objective lenses.

Only microscope oil and ethanol are used. Exposure to the former is controlled by wearing gloves (see COSSH form)

Refer to individual biological risk assessment (please specify)
Ditto
	Yes

Yes
	Low

Low


	(Class3B only - check whether this applies to your system)

	Environment & People:

Small room - Risk of tripping or inadvertent damage to system due to overcrowding.

Use by inadequately trained users

Prolonged use of equipment can result in eye strain and ergonomic hazards
	Users 

Users

Users
	The room is equipped and intended for a maximum of 5 people in normal circumstances, as explained during training.

All users must be authorised; before this can occur they must have received proper training. They must also sign a declaration that they have read the local rules.  

Ambient lighting kept as high as possible, restrict length of time using equipment, workstation designed for ease of use and fully adjustable chairs provided.
	Yes

Yes 

Yes
	Low

Low

Low
	





















IR075

RA_Confocals_example

2

